## Slide 1: In-ear EEG Auditory Steady State Response Analysis
Integrated Systems Laboratory (ETH Zürich)
Benedetta Mazzoni b.mazzoni@unibo.it
Andrea Cossettini cossettini.andrea@iis.ee.ethz.ch
Simone Benatti simone.benatti@unimore.it

## Slide 2: Setup description (1)
Stimuli (for ASSR)
  Over-the-head or bone-conduction headphones
Electrodes
  earEEG electrodes: fabricated by Dätwyler Schweiz AG
Naison
headphones
Ear-EEG electrode fabricated by Dätwyler Schweiz AG and connected to an active buffering PCB
Electrodes configuration on a test subject
Channel | Location | Type
CH1 | In-ear (left) | Ear-EEG electrode + active buffering
CH2 | In-ear (right) | Ear-EEG electrode + active buffering
CH3 | front of the ear (left) | Wet electrode + active buffering
CH4 | back of the ear (left) | Wet electrode + active buffering
CH5 | front of the ear (right) | Wet electrode + active buffering
CH6 | back of the ear (right) | Wet electrode + active buffering
REF | Mastoid (left) | Wet electrode + active buffering
BIAS | Mastoid (right) | Wet electrode

## Slide 3: Setup description (2)
Data acquisition platform
  BioGAP (https://arxiv.org/abs/2307.01619)
    PULP-based computing platform (GAP9)
    TOPS capabilities, 10 RISC-V cores, state-of-the-art performance for tinyML
    Nordic nRF52 for BLE connectivity
    Can be flexibly connected to a large variety of sensor interfaces
    ADC for biopotentials (ADS1298, 8 ExG channels, 24 bit resolution)

## Slide 4: Validation measurements: Alpha waves
Protocol
  Subject alternating eyes open and eyes closed with 10-second intervals
Goal
  Analyze the power of alpha waves within the frequency range of 7 to 13 Hz
Data processing
  Filtering: band-pass, Butterworth (order 10, cutoff: f1=0.5 Hz, f2= 40 Hz)
  Filtering: notch, order 2, at 50 Hz
  Spectrogram computation: 1024-samples windows (2 seconds) with 768-samples (1.5 seconds) overlap

## Slide 5: Validation measurements: Alpha waves
Results (Subject A)
  Left channels
  Alpha activity visible in the 7-13 Hz range
Channel 1 in-ear
Channel 3
Channel 4
eyes closed
eyes open
eyes open
eyes closed
eyes open
eyes closed

## Slide 6: Validation measurements: Alpha waves
Results (Subject A)
  Right channels
  Alpha activity visible in the 7-13 Hz range
Channel 2 in-ear
Channel 5
Channel 6
eyes closed
eyes open
eyes open
eyes closed
eyes open
eyes closed

## Slide 7: Validation measurements: Alpha waves
Results (Subject B)
  Left channels
  Alpha activity visible in the 7-13 Hz range
Channel 1 in-ear
Channel 3
Channel 4
eyes closed
eyes open
eyes open
eyes closed
eyes open
eyes closed

## Slide 8: Validation measurements: Alpha waves
Results (Subject B)
  Right channels
  Alpha activity visible in the 7-13 Hz range
Channel 2 in-ear
Channel 5
Channel 6
eyes closed
eyes open
eyes open
eyes closed
eyes open
eyes closed

## Slide 9: ASSR: protocol
Protocol description
  5 healthy volunteers
  Subjects wear the electrodes and the headphones. During acquisitions, they are allowed to read
  Auditory stimulus from the headphones:
    sinusoidal tone with a carrier at 4 kHz and AM modulated [1][2] (we test 80 Hz and 88 Hz)
    Intensity: 65 dB SPL (sound pressure level)
    Duration: 15 minutes (similar to [3])
  BioGAP sampling frequency: 500 Hz
    [1] Kubota, Toshinori, et al. "Detecting the recruitment phenomenon in adults using 80-Hz auditory steady-state response." Auris Nasus Larynx 46.5 (2019)
    [2] Herdman AT, Lins O, Van Roon P, Stapells DR, Scherg M, Picton TW. Intracerebral sources of human auditory steady-state responses. Brain Topogr. 2002.
    [3] C. B. Christensen, T. Lunner, J. M. Harte, M. L. Rank and P. Kidmose, "Chirp-Evoked Auditory Steady-State Response: The Effect of Repetition Rate," in IEEE Transactions on Biomedical Engineering, vol. 69, no. 2, pp. 689-699, Feb. 2022.

## Slide 10: ASSR: data processing
Signal processing pipeline
  Filtering: band-pass, Butterworth, order = 10, cutoff 30 – 180 Hz
  Filtering: notch, at 50 Hz (PLI) and 100 Hz (2nd harmonic of PLI): order = 2, quality factor = 100
  Power Spectral Density (PSD) analysis using WELCH Method:
    Windowing = 5 sec
    Discrete Fourier Transform (DFT) of the periodgram
    Squared magnitude, yelding power spectrum estimates for each segment
    Averaging of the individual power spectra (to reduce noise floor)
  Result: array of power measurements vs. frequency «bin» (Power/Freq) estimated at frequency intervals.
  Automatic peak detection (peak of the response at the AM frequency 80 Hz)
  Automatic Signal-to-Noise Ratio (SNR) estimation

## Slide 11: Results: 80 Hz AM
Power spectra
  Signal strongly visible in all channels

## Slide 12: Results: 80 Hz vs 88 Hz AM stimuli
Power spectra comparison
  Signal strongly visible in both cases (in-ear channels)

## Slide 13: Results: over-the-ear vs bone-conduction headphones
Power spectra comparison
  Signal strongly visible in both cases (in-ear channels)